HiFliQ Protein G FPLC Columns

Datasheet

Protein Ark’s HiFliQ Protein G FPLC columns pre-packed and ready to use with Protein G Agarose FF resin for simple, one-step and rapid antibody purification.

The Agarose is coupled with recombinant Protein G at 3.5 mg/ml of resin providing high affinity and high capacity immunoglobulin binding via the heavy chain of the FC region over a broad range of source species and subclasses (see user guide for further details).  Compatible with all common chromatography instruments (including ÄKTA™ FPLC’s).

| Resin Specification: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item: | HiFliQ10-PG-1 |  | HiFliQ20-PG-1 |  | HiFliQ50-PG-1 |
| Column Volume: | 10 ml |  | 20 ml |  | 50 ml |
| Resin: | Protein G Agarose FF |  | Protein G Agarose FF |  | Protein G Agarose FF |
| Typical Binding Capacity: | 200 mg/ml |  | 400 g/ml |  | 1000 mg/ml |
| Base Matrix: | Agarose |  | Agarose |  | Agarose |
| Mean Bead Size: | 45-165 µm |  | 45-165 µm |  | 45-165 µm |
| Recommended flow rate: | 10 ml/min |  | 10-20 ml/min |  | 10-50 ml/min |
| Max. operating pressure: | 0.5 MPa (72 psi) |  | 0.5 MPa (72 psi) |  | 0.5 MPa (72 psi) |
| External Dimensions: | 16mm ID x 50mm BH |  | 16mm ID x 100mm BH |  | 26mm ID x 100mm BH |
| Column Construction: | Acrylic |  | Acrylic |  | Acrylic |
| Top and bottom filters: | Polyamide |  | Polyamide |  | Polyamide |
| Top and bottom plugs: | Polypropylene |  | Polypropylene |  | Polypropylene |
| Inlet Port: | 10-32 (1/16”) Female |  | 10-32 (1/16”) Female |  | 10-32 (1/16”) Female |
| Outlet Port: | 10-32 (1/16”) Female |  | 10-32 (1/16”) Female |  | 10-32 (1/16”) Female |
| Storage (2-8°C): | 20% ethanol |  | 20% ethanol |  | 20% ethanol |

| Ordering Information: |  |  |
| --- | --- | --- |
| Product | Units | Order Code |
| 10 ml HiFliQ Protein G FPLC column (1 x 10 ml) | 1 x 10ml | HiFliQ10-PG-1 |
| 20 ml HiFliQ Protein G FPLC column (1 x 20 ml) | 1 x 20 ml | HiFliQ20-PG-1 |
| 50 ml HiFliQ Protein G FPLC columns (1 x 50 ml) | 1 x 50 ml | HiFliQ50-PG-1 |
| *Custom sizes also available |  |  |

Instruction of Use:

Each packed column is sealed with a pressured syringe in the bottom end of the column.

It is then placed in sealed plastic bag.

1. Cut the plastic bag to take out the packed column with care.

2. Follow the flow direction to clamp the column to a vertical stand.

3. Get the chromatography system ready for connection.

4. Unscrew the stop plug. Some liquid may come out.

5. Connect the column top to the inlet of the chromatography system.

6. Gently unhook the springs from the shaft top of the bottom syringe using balanced

force.

7. Twist with push-up force to unscrew the male-thread adaptor from the column top.

Keep this storage syringe for later use.

8. Connect the bottom side of the column to the chromatography system.

9. Run at reduced flowrates to wash away the storage solution then run at normal

flowrate until the column is equilibrated.

DO NOT EXCEED 5 BAR

Storage after use:

For short term storage, seal both ends of the column with stop plugs if it isn’t being used.

For longer term storage, follow the instructions below to store the column.

1. Fill the column with suitable storage solution. Typically it requires to pass at least 2

CV storage solution.

2. Stop the pump.

3. Suck the storage solution to fully fill the storage syringe (including the male-thread

and the tubing part). Leave the syringe upside-down and push out the air bubble.

Adjust the liquid level inside the syringe to 3.5 ml to 4 ml (for 0.8 to 1.1 cm ID

columns) or 11 to 12 ml (for 2.6 cm ID columns).

4. Carefully screw the male thread part of the syringe system to the outlet of the

column. Finger tight is enough.

5. Dis-connect the column top from the chromatography system.

6. Screw a 1/16” stop plug to seal the column top.

7. Use balanced force to hook the springs back the top shaft of the syringe.

8. Check and make sure no leakage.

9. Place the column in a cold room.